Proteonano™ HT Quality Statement (PQS)

Version: V.1.0

Date: 2026-03-20

Quality Control Summary for Samples and Plates

General description of quality control methods and performance outcomes for the Proteonano™ HT Plasma Proteomics Assay

## 1. Document Overview

Platform
Proteonano™ HT Assay

Calibration and QC System
PACS – Proteonano Assay Calibration & Standards

Application
High-throughput, large-cohort plasma proteomics

Client

AAA

StudyID

9c4ad311-3876-4eb0-ad9f-eb88e91969b4

Date

2026-03-20

Intended Use
Research Use Only (RUO)

### Document Purpose

This Proteonano™ HT Quality Statement (PQS) summarizes the quality control, calibration, normalization, and performance evaluation applied to samples processed using the Proteonano™ HT assay. The purpose of this report is to provide transparent, standardized documentation of assay performance across samples, plates, and analytical runs, supporting confidence in data quality, reproducibility, and cross-batch comparability.

## 2. Quality Control Framework Overview

The Proteonano™ HT assay incorporates a multi-layered quality control framework embedded directly into the 96-well workflow. Quality controls are strategically placed to monitor distinct sources of technical variability arising during sample preparation, plate-based processing, and LC–MS analysis.

Embedded QC Architecture

| QC Level | Control | Purpose |
| --- | --- | --- |
| Instrument | HEavy-Labeled Peptides (HELP) injections | LC–MS response calibration |
| Plate | Pooled plasma QC with HELP | Plate effect monitoring |
| Sample | Sample with HELP | Sample-specific calibration |
| Study | Historical QC | Cross-batch alignment |
| Blank | Matrix-free blanks | Contamination / carryover |

QC elements are embedded at their point of origin, enabling targeted correction rather than global post hoc normalization.

## 3. Calibration & Normalization Summary

### Sample Summary

The table below summarizes the total number of biological samples and associated controls evaluated in this study, including pass and flag counts based on established quality criteria. Samples flagged by quality metrics are evaluated on a case-by-case basis to assess the appropriateness of normalization and their impact on downstream quantitative analyses.

Table: Sample & QC Summary

| Category | Total | Pass | Flag |
| --- | --- | --- | --- |
| Internal Calibrator | 36 | 36 | 0 |
| QC samples | 42 | 42 | 0 |
| Samples | 1038 | 1038 | 0 |
| Blanks | 0 | 0 | 0 |
| Sum | 1116 | 1116 | 0 |

### Quality Control Metrics & Flags

The table below summarizes the quality control metrics evaluated for the Proteonano™ HT assay, along with their corresponding acceptance criteria and assessment outcomes. These metrics are designed to monitor assay performance across sample preparation, plate-based processing, and LC–MS detection. Each metric is evaluated independently, and results are reported as pass or flag based on predefined thresholds. Metrics that fall outside acceptance criteria are flagged for review rather than automatically excluded, allowing informed decisions based on normalization performance and study context.

Table: QC Metrics and Acceptance Criteria

| Metric | Evaluated On | Acceptance Criteria | Result |
| --- | --- | --- | --- |
| Instrument factor | HELP injections | -0.25~0.60 |  |
| Plate factor | Pooled QC | -0.70~0.70 |  |
| Sample factor | Samples | -0.80~1.80 |  |
| Median factor | Samples | -2.15~1.40 |  |
| Study factor | Inter-Study QC | -0.17~0.17 |  |

### Plate-Level Normalization

Plate Scaling is derived from the median of reference signal ratios calculated from pooled plasma QC samples distributed across each plate, yielding a single plate-level scaling factor applied uniformly to all samples on the plate. This scaling primarily reflects systematic variability introduced during plate-based processing, including liquid handling performance. The plate scaling factor acceptance criteria is defined as {{plate_scaling_acceptance_range:-0.64-0.66}} per plate.

| Plate ID | Scale Factor | Acceptance Range | Status |
| --- | --- | --- | --- |
| 1 | 0.163867 | -0.70~0.70 | PASS |
| 2 | -0.003657 | -0.70~0.70 | PASS |
| 3 | 0.004877 | -0.70~0.70 | PASS |
| 4 | -0.1717 | -0.70~0.70 | PASS |
| 5 | 0.456499 | -0.70~0.70 | PASS |
| 6 | 0.089672 | -0.70~0.70 | PASS |
| 7 | -0.637791 | -0.70~0.70 | PASS |
| 8 | -0.273208 | -0.70~0.70 | PASS |
| 9 | 0.001935 | -0.70~0.70 | PASS |
| 10 | 0.655579 | -0.70~0.70 | PASS |
| 11 | -0.180723 | -0.70~0.70 | PASS |
| 12 | 0.454103 | -0.70~0.70 | PASS |
| 13 | -0.263246 | -0.70~0.70 | PASS |
| 14 | 0.371854 | -0.70~0.70 | PASS |
| 15 | -0.067127 | -0.70~0.70 | PASS |

Plate scaling factors were within predefined limits, indicating stable plate-based processing.

### Sample-Level Calibration & Normalization

Sample-level calibration factors were evaluated across all study samples to assess normalization performance and signal consistency. Individual sample scale factors are summarized using distribution-based metrics and pass/flag counts rather than reported individually. This approach provides a robust overview of sample-level performance while avoiding over-interpretation of isolated sample-level variability.

| Metric | Result |
| --- | --- |
| Total samples evaluated | 1038 |
| Samples within acceptance range | 1038 |
| Flagged samples | 0 |
| Sample scale factor range | -0.80~1.80 |
| Median sample scale factor | ~0.00 |